聚乳酸系骨釘骨板之短期活體內組織反應與降解變化

Tissue response and biodegradation of PLA bone screw and plate

中文摘要

到目前為止，使用聚乳酸植體最大的限制仍是其強度不足，故多用於海綿骨骨折之固定，製備時可以緩慢降溫方式來提高聚乳酸結晶度，增加聚乳酸的強度，但缺少此種聚乳酸的組織學研究，且有關聚乳酸之研究多為長期觀察，缺乏對於聚乳酸植體在海綿骨骨折癒合關鍵的前五週時期之組織學及物理性質變化的探討。所以本研究的目的在探討以緩慢降溫方式所製備之聚乳酸poly (L-lactide) 在體內35天期間的組織反應及物理性質變化。使用日本島津公司重量分子量21萬之聚乳酸，以220 ℃熱熔法，在氮氣中自220 ℃緩慢降溫產生結晶度43 %的聚乳酸，並切割成2 x 6 x 25 mm3試樣，經環氧乙烯滅菌後，植入在Sprague Dawley大鼠背部皮下組織，經1天、3天、7天、21天、35天五個試驗期間，以石臘及樹脂標本之組織學觀察，評估其組織反應並以三點彎曲強度、黏度及熱性質檢測，評估其物理性質變化。滅菌前聚乳酸植體試樣三點彎曲強度為120 MPa左右，結晶度為43 %左右，經55 ℃、2.5小時環氧乙烯滅菌後，發現滅菌前後之物理性質沒有差異。就聚乳酸植體試樣組織癒合過程而言，可分為止血期、發炎期、增生期及重塑期四個時期。術後第一至三天為止血期、發炎期，是手術外傷所造成之發炎反應，炎症細胞以中性白血球為主，術後第三天，炎症細胞以單核球細胞為主，可見纖維母細胞分布。術後第七天為增生期，炎症細胞分布範圍變小，可見單核球細胞數目減少，此時可見含有許多增生纖維母細胞及血管的纖維包膜形成，由樹脂磨片標本可見纖維包膜與聚乳酸植體緊密接觸。植入後第21天為重塑期，可見纖維包膜的厚度比第7天薄，包膜內之纖維母細胞變較細長，血管數目減少。植入後第35天亦屬重塑期，纖維包膜厚度比第21天薄，包膜內之纖維母細胞及血管數目減少。所以此聚乳酸植體試樣之組織反應是以單核球細胞為主要的炎症細胞反應，植入後第七天形成纖維包膜，且隨植入時增加，纖維包膜愈成熟，在植入35天期間未見多核吞噬細胞。就聚乳酸植體試樣物理性質變化而言，經35天試驗期間其分子量沒有明顯差異，故沒有明顯降解反應發生，能維持一穩定之三點彎曲強度、黏度及熱性質，所以植入周圍組織對此聚乳酸植體試樣反應溫和，纖維包膜隨植入時間增長而愈成熟。這可能是因為具有足夠高之初始分子量且在製備時置於氮氣中緩慢降溫，聚乳酸主鏈未受製備過程產生大幅熱降解而製備時不予施壓致使材料具有適當之結晶度，雖然進一步長期降解研究是必須的，此種製備方式所得之聚乳酸植體為一可維持五週機械強度之生物鈍性材料。

英文摘要

The purpose of this inverstigation is to study the tissue response and degradation of PLLA in bone tissue. The materials used in this study included PLLA bone screws and bars which were composed of 5% D-form and 95% L-form polylactide. 18 PLLA bone screws with 8 mm in length, 2mm in diameter,24 PLLA bars (2×6×2.5 mm3)and 6 PLA plate were studied and 6 commercial plate and screws were control, (manufactured by Macropore) were control. 24 New Zealand white rabbits (3-4 months old, 3-4 kg weight) were the test animals, and they were divided into three groups (A and B、C) proceeding in this study. In group A：One PLLA screw and one commercial screw were implanted in the right tibia. Two PLLA screws were implanted in the left. Furthermore, one PLLA bar was implanted in both right and left tibia respectively with onlay model in group B. Also,one PLLA plate was implanted in the left tibia and one commercial plate in the right respectively with onlay model in group C . Many tests were done on each sample in 1, 4, 8 and 12 weeks which included the observation of tissue response in group A, and the change of three-points bending, weight loss, molecular weight, crystallinity and morphology of fracture surface by scanning electronic microscopy（SEM）in group B, and the change of three-points bending, weight loss and molecular weight in group C. It was clear that the tissue response in group A revealed good biocompatibility.Otherwise, it was showed that there was no difference in weight loss until 12 weeks, and the initial bending strength was 147-157 MPa. The molecular weight also decreased from 102 kDa to 67 kDa. It is clear that the bending strength decreased while the molecular weight decreased during the degradation process, but there was no weight loss at the same time in group B. In group C：The initial bending strength of commercial plate is stronger than controle（50 MPa：35 MPa）which is due to its thickness.Though, that strength decreased to 78% in 4 weeks and 58% in 12 weeks compared with controle group：100% in 4 weeks and 74% in 12 weeks. It is evident that the self-manufactured plate by injection method has lower molecular weight but maintain strength long enough than commercial one. Conclusion： These results suggest that the BTO screws and plates are good biocompatable. They are potential products in the field of oral — maxillofacial devices in the future.